Table of Contents

1.0   Introduction………………………………………………………….…………… 2

1.1  Motivations

1.2  General Description

1.3  Features

2.0  Specifications..............................…..................………........…………..…….........  5

2.1  System Specifications

2.2  Mechanical Specifications

2.3  Electronic Specifications

3.0   Detector assembly and installation..............………………….…………………. 8

3.1  Mounting NERO blocks.

3.2  Placing NERO proportional counters

3.3  NERO logic

3.3.1  Trigger system

3.3.2  NERO electronic diagram

3.4  NERO connections

3.4.1  Counter connections

3.4.2  TDC connections

3.5  Biasing the detectors

4.0   DAQ.……………………………………………………….….………………….. 14

4.1  Starting the readout

4.2  Starting specTcl

4.3  Starting the scalers display

4.4  How the spectra should look like

5.0   Electronic control.………………………………………………….……………. 17

5.1  Shapers: gains and thresholds

5.2  ADC thresholds

5.3  Electronic checks

6.0   Operation .........................….…….........................…........…....……….....….……19

6.1  Background

6.2  Efficiency

7. 0  Principle of the measurement……………………………….………….…….…. 21

7.1  Pn value measurement

##### INTRODUCTION

  - Motivations

The r-process is a major mechanism for producing elements heavier than Fe. It occurs in an environment of large neutron flux, where nuclei experience neutron capture faster than the competing (-decay. Within a given isotopic chain, the neutron captures compete with the inverse photo-neutron disintegration until equilibrium is established between both reactions. At this point, an equilibrium isotope distribution is built up, with maximum abundances found in the so-called waiting point nuclei. After a while, these nuclei decay by (-emission, giving rise to new heavier elements. The process is repeated sequentially at each isotopic chain, so that the flux of r-process nuclei proceeds through the neutron-rich side, very close to the neutron-drip line. When the neutron flux diminishes, the exotic nuclei will decay to stability by (-decay and (-delayed neutron emission (freeze-out). As a consequence of delayed neutron emission decay, a fraction of decay products move from isobaric line A to line A-1 during the freeze-out, thus converting the freeze-out abundance pattern into the final r-process abundance pattern. Fig.1 illustrates the main effect of the (-delayed neutron emission during the r-process. As can be seen, the inclusion of this process in r-process calculations smoothes the calculated final abundance of element as in Fig.1.

Fig.1: Smoothing effect of the delayed neutron emission on the r-process model calculation. Because the (-decay does not change A, the difference in the abundance pattern, before and after the freeze-out, are due to the (-delayed neutron emission.

In this context, the knowledge of the probability for (-delayed neutron emission (Pn value) is a critical input for r-process models.

The NERO detector (Neutron Emission Ratio Observer) was designed and built at NSCL in order measure the (-delayed neutron emission probabilities.

  - General description

NERO was mainly conceived to work in combination with the NSCL Beta Counting System (BCS). It consists of three concentric rings of proportional counters (one inner 3He and two external BF3 gas tubes rings) embedded in a 60 x 60 x 80 cm3 polyethylene matrix, with a 22.4 cm diameter central cylindrical hole to hold the Beta Counting System (BCS) inside (see Fig.2). The moderator system was motivated by the necessity of using the slow-neutron detection technique, which highly maximizes the detection efficiencies (see Knoll’s book, chapter 14). In order to make a compact detector, easily manageable and fashionable, a polyethylene moderator matrix was chosen because its higher density of protons compared with other existing materials with higher moderation efficiencies.

The neutron detectors chosen were proportional counters filled with 3He or BF3 gas. In such detectors, neutrons are indirectly measured from the charged particles released in the reactions 3He(n,p)3H and 10B(n,()7Li. In both cases the incoming neutron energy is very small compared with the Q-value of the reactions so that the energy of the reaction products is essentially the Q-value itself. If one of the two charged particles is not stopped entirely in the counter, the amount of energy recorded is smaller than the one carried by the neutron. This effect, called wall effect, determines the shape of the neutron spectra (see Knoll chapter 14).

#### Fig.2: View of NERO with all the proportional counters in place and connected to the preamplifiers.

In the final setup, layers of boron carbide and water can be placed around the detector to minimize neutron background mainly beam-induced and from cosmic rays.

  - NERO Features

This is a list of the main advantages of the NERO design:

- The moderating polyethylene material has a relatively high density of light nuclei (carbon and hydrogen), which are optimum to moderate neutrons.

- The polyethylene matrix has a simple assembly scheme making the whole detector easily portable.

- The 4( coverage angle allows for a high geometrical efficiency.

- Because of the use of long counters and ring structure, NERO has a rather flat efficiency response as a function of the energy (see Fig.3)

- The special configuration of NERO–with a central hole in the center of the polyethylene matrix–leaves room to insert a segmented implantation (-decay detector (e.g. the Beta Counting System at NSCL), allowing for detection of (-delayed neutrons (see Fig.2).

Fig.3: NERO efficiency vs. neutron energy as calculated with the MCNP code during the design phase. The plot includes the total NERO efficiency(empty circles), as well as the efficiency of the each of the three rings of proportional tubes.  Ring 1 (squares) is the inner ring, Ring 2 (empty diamonds) is the middle ring and Ring 3 (empty triangles) is the outer ring.

##### SPECIFICATIONS

  - System specifications

- Detectors type: one inner ring with 16 3He proportional counters, and one middle and outer ring with 20 and 24 BF3 proportional counters.

- Measuring energy range: neutrons of energy from 1keV to 5MeV.

- Efficiency: 45% relatively constant from 1keV to 500keV, dropping off to around 26% at 5MeV (values from simulations).

  - Mechanical specifications

- Length: 80cm

- Width: 60cm

- Height: 60cm

- Weight: 400 pounds

- Internal beam line opening diameter: 22.4 cm

- Inner ring radius: 13.6 cm

- Second ring radius: 19.2cm

- Outer ring radius: 24.8cm

2.3   Electronic specifications

- Proportional counters high power:

| Quadrant | 3He proportional- counters voltage (V) | BF3 proportional -counters voltage (V) |
| --- | --- | --- |
| A | 1130 | 2620 |
| B | 1130 | 2620 |
| C | 1084 | 2620 |
| D | 1350 | 2620 |

Fig.4: Diagram of NERO quadrant and detector labeling. The beam is coming out of the page.

- Electronic modules:

- CAEN VME ADC (model V785)

- CAEN VME TDC (model V767)

- SIS Scaler (model SIS3820)

- Pico System CAMAC shaper-discriminator

- Preamplifier signals: analogic positive with amplitude from 20-30mV (see Fig.5)

- Pico System shaper-discriminator (see Fig. 6) output:

- Shaper output signals (going to ADC input): Negative

- Discriminator (TDC input): Negative standard ECL

- Logical OR of all the 16 discriminator signal (MG): Negative standard NIM

- One channel shaper test

- One channel discriminator test

- Discriminator channels sum: provide a multiplicity signal of 50mV for channel hit.

Fig.5: Typical signals from NERO preamplifier 3He (left) and BF3 proportional counters (right).

Fig.6: Front view of two Pico System shaper-discriminators corresponding to quadrant A (left) and B (right).

##### NERO ASSEMBLY AND INSTALLATION

# Mounting NERO blocks

NERO is mounted on a special support that allows a fine adjustment of the position and alignment with the beam-line. It sits on two layers of four rectangular polyethylene blocks. The bottom layer is labeled with the letter “B” (bottom) and is marked with an arrow in the direction of the beam, and a circle indicating how the four pieces go together. The next layer, marked with the letter “T” (top), is assembled in the same way. Several thin plastic sheets are also included in the assembly setup in order to make the intersection point as smooth as possible before adding the blocks holding the tubes. If these two base layers are not as flat as possible, it can be difficult to insert the detector tubes inside NERO. The blocks that hold the proportional counter tubes are labeled with B1…B6 and T1…T6.  The six “B” blocks are set on the two base layers, with the block labeled “B1” positioned upstream and the arrow pointing in the direction of the beam. The blocks labeled “T” are set on the lower sections, again with “T1” positioned upstream and the arrow in the direction of the beam.  Eight threaded rods run through the blocks and are tightened with washers and nuts at each end.  Because of difficulty aligning the blocks after they are all installed, it is recommended that the threaded rods be pushed though each section as it is installed on the stack. The nuts on the ends of the rods should be left slightly loose until all the detector tubes have been installed, as they require the blocks to move somewhat in order to go in easily.

  - Placing NERO proportional counters

Before inserting the BF3 proportional counters make sure that the connections are tight, and that the screw in the end cap of the detector is tight as well (see Fig.7).

NERO is divided into four quadrants and the detectors are numbered according to quadrant (see Fig.4). The 3He proportional counters of each quadrant need to be powered at a different voltage, so keeping track of the detector positions is important.

Fig.7:  BF3 assembly

  - NERO logic

    - Trigger system

The module used as trigger is the CAEN VME V262. It has four inputs (IN0-3) and eight outputs (Shp0-3 and OUT 0-3).  IN1-3 and OUT1-3 are used as control signals to be read/set with the Readout software.

| Plug name | Signal meaning |
| --- | --- |
| IN0 | Master Gate Live |
| Shp0 | Computer Going Software Busy (not used) |
| Shp1 | Computer Going Ready |
| Shp2 | Trigger Acknowledge |
| Shp3 | Module Clears |

These signals should be hooked to the experimental electronics as shown in Fig.8:

Fig.8: VME Trigger electronics block diagram

The Master Gate Live trigger (MG Live) comes in from the experiment trigger logic to the first trigger latch module, which produces a trigger signal going to IN0. This signal is held asserted as long as the computer is required to react to it. When the computer reacts to the trigger, it pulses SHP2 (Trigger Acknowledge), de-asserting the trigger latch module.

The Master Gate Live Trigger also sets the second busy latch module, which inhibits further gates from the experimental electronics. When the computer has finished reading an event it pulses SHP3 (Module Clears), which can be used to clear modules that require external clear gates (typically CAMAC standard modules). Finally, it pulses SHP1 (Computer Going Ready) to clear the busy latch module, enabling the electronics to gate the next event into the computer.

3.3.2. NERO electronic diagram

Fig.9: NERO logic diagram.

In the NERO standalone version, a Master Gate signal (generated from a detected neutron), in coincidence with the non-busy signal from the latch module 2, generates the Master Gate live, which sets busy all the latch modules and generates the ADCs gate. The latch module 1 maintains the Master Gate Live signal (going to the V262 IN0) true until the computer responds to it (sending the Trigger Acknowledge signal from SHP2). The busy output signal from the latch module 3 generates the 200 (sec TDC Start gate (going to the V262 IN2 channel), while the non-busy signal (time-out condition provided by the 200 (sec delayed Master Gate) informs NERO that the TDC gate is closed and that the readout of the TDC is ready for the next event (see section 4 for details about this condition). The latch module 4 is used to set a condition to verify that data are available to be read. After the readout of all modules is done, the V262 SHP3 signal provides the Clear Gate for all possible CAMAC modules (not used in the current setup) and the latch module 4. Finally, a Computer Going Ready signal, coming from the V262 SHP1, unlatches the latch module 2, providing a new non-busy signal for the coincidence with the Master Gate.

Shaper and discriminator signals are the inputs for the ADCs and the TDC, respectively.

Note that in the experimental version, the Master Gate signal is provided by a (-decay (electron) detected by the Beta Counting System.

  - NERO connections

Most connections are straightforward. We just point out here those that can be problematic.

    - Proportional counters connections

NERO proportional counters are connected to dedicated preamplifiers via HV cables, either RG-58 (for B3F) or RG-158 (for 3He). The preamps can be a source of noise and sparks. The first remedy for this is cleaning them inside, and making sure they are correctly grounded; if the noise or sparks persist, then bring them to the technician.

There are two different types of connections for the 3He tubes: one is a basic standard RG-58 HV connection and is self-explanatory; the other type features two threaded connectors; one of these threaded connectors has a small metal tab attached, indicating the correct connector to be used (see Fig.10).

Fig 10:  3He proportional counter. The picture shows threaded connectors that have to be used for the connection.

    - TDC connections

The CAEN TDC VME module V767 is used in the NERO setup because of its capability to detect multiple events. In the START GATING mode, a start input signal opens a gate during which multiple stop signals can be detected. These stop signals are provided by each 16-channel discriminator output (four modules = (16 - 1) x 4 = 60 detectors), which were produced by the detection of a neutron with any of the sixty proportional counters.

The front side of the NERO TDC shows a 8-pin input connector defining the particular used mode. In the START GATING mode (used in the NERO setup) it is not possible to connect the start signal with an 8-flat ribbon channel cable, because the TDC will expect additional inputs, thus creating ambiguity in the mode selection. Therefore, a single double twisted cable, coming from the latch module 3 is used as START signal (the right connection orientation is indicated with the word “UP” as seen Fig.11). The multiple stop signals arriving from the four discriminators can be detected during the time the gate is opened.

It is important to emphasize that the TDC module accepts positive polarity signals only; so, the cables coming from the discriminators have to be twisted. This does not mess up channels because the enumeration of the channels for the TDC starts from the bottom while the discriminator one from the top.

Fig.11: TDC connections. Note the single twisted cable used for the TDC gate with the word “UP” indicating the right orientation.

3.5    Biasing the proportional counters

ALWAYS BE VERY CAREFUL WITH THE HV! DON'T TOUCH PREAMPLIFIERS OR CABLES GOING FROM PREAMPLIFIERS TO TUBES WHEN BIASED!

NERO needs special modules for the high voltage supply. Two of them fill the full width of the rack (MODELS?); they provide the HV for the B3F tubes. The working voltage is 2620V. Four single-slot modules (MODELS?) are used for the He3 proportional counters for the four quadrants. The values that must be used to bias the detectors are listed in the electronic specifications (section 2.3).
     Make sure you don't turn the knobs below zero when biasing the detectors down. This is possible in some cases. If this happens the dial scale will be misaligned and the voltage will be too high the next time you bias up. The tubes could be fried! Make sure the detectors are bias both up and down slowly.

##### DAQ

4.1   Readout

Before you can display any of the spectra, you need to run the readout program that grabs the data from the different VME modules. To start the software, you must be logged in to spdaq28.nscl.msu.edu as user NEROtest. Go to directory NEROtest/Standalone/readout and type './Readout'. The program is started in the shell typing 'begin'.To stop it, type 'end'; to exit the program type 'exit'.

The data are fed into three different modules: Scalers, TDC and ADCs. The scalers use a specific scaler display while the ADCs and TDC are displayed with SpecTcl.

The readout can also be started with a special tcl script, which is highly recommended. In the directory NEROtest/Standalone/readout type './nero_guiproRead' to get the tk screen and click on BEGIN.

Once the readout is running, it is possible to display the data buffer by typing ‘./nero_bufdump’ which runs the tcl script NERObufdump.tcl.

4.2    SpecTcl

The program Spectcl is located in the directory NEROtest/Standalone/spectl and to run it you have to type ./Spectcl .

Once the Xamine, SpecTcl, SpecTcl Graphical User Interface, and TkCon 1.6 SpecTcl windows have appeared, press the Clear Spectrum button followed by the Attach Online button (in the SpecTcl Window). This will allow you to look at the online data.

Spectra can be seen in the Xamine Window by clicking on Window, followed by Read Configuration. Select the Windows directory and choose the desired configuration file.

4.3    Scalers
     The scalers count every event that is above the discriminator thresholds. The display shows the rate and total number of events of the current run. The scalers are used for efficiency measurements, background control and NERO performance. The scaler display can be started typing './NEROgoscalers' in directory NEROtest/Standalone/scalers.

4.4    How the spectra should look like
     The energy spectra (see Fig.12) do not represent the energy of the neutrons, but are characteristic of the proportional counters. In both reactions indeed, the incoming neutron energy is very small compared with the Q-value of the reaction, so the energy of the reaction products is essentially the Q-value itself.

The shape of the spectra is determined by wall-effects (see Knoll chapter 14).

The energy spectra have three contributions: low energy noise-signals that include ADC pedestals, neutron signals, and background. The noise forms a large and shaped peak, which can be discriminated from the real signal by setting the discriminator thresholds (and ADC thresholds for pedestals). The size of the noise-peak is generally so large that, if present, renormalizes the display in such a way that the real neutron signals can only be seen in logarithmic scale.

Since the “good” neutron signals and neutron background signals have the cover a similar range of energy and have very similar shape, it is impossible to separate them by software or hardware conditions. This requires running long background runs in order to subtract the background rates from the total measured neutron rates.

The TDC spectra for the standalone version are shown in Fig.13. Because the signal used as STOP is the delayed START, all the counts appear roughly at the same time, and the spectra looks like a big peak plus background and noise.

Fig.12: ADC spectra collected with a 252Cf 50 (Ci source for a five minutes measurement. The first four spectra correspond to the 3He counters from quadrant C. The vertical straight lines identify the thresholds used for the discriminator. Looking at these spectra is indeed the way to set such a thresholds..

Fig.13: Standalone TDC spectra from quadrant B, collected with a 252Cf source during 5 minutes. Channels 2 and 10 have been expanded to show the main peak produced by the constant time delay between the start and delayed start (or equivalently stop) signals (see text for details). The other part of the spectra is noise and background. The first four spectra correspond to the He3 proportional counters.

##### ELECTRONIC CONTROLS

# 5.1   NERO Shaper/Discriminator gains and thresholds

Shaper gains and discriminator thresholds are controlled via software from the Readout program.

Shaper gains are set in file: NEROtest/Standalone/readout/shaper_init_NERO.dat. The maximum gain is 255; if a greater number is used, the value of the gain is set to the difference to 255.

Discriminator thresholds are set in file: NEROtest/Standalone/readout/shaper_init_NERO.dat. The maximum threshold level is 255, if a greater number is used the value of the gain is set to the difference to 255.

Fig.14 shows an example of shaper and discriminator file. The crate, branch, slot number and channel set as test channel are specified for each quadrant after the corresponding text lines. The other numbers define the gains (for shapers) and the thresholds (for discriminators).

NERO_QUADA_Shaper 0 3 2 1 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 NERO_QUADB_Shaper 0 3 4 1 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 NERO_QUADC_Shaper 0 3 6 8 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 NERO_QUADD_Shaper 0 3 8 5 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255 255

NERO_QUADA_CFD 0 3 3 1 35 25 15 25 25 25 25 25 25 25 25 25 25 25 25 25 NERO_QUADB_CFD 0 3 5 15 25 20 23 21 20 24 21 22 23 18 32 22 20 20 40 255 NERO_QUADC_CFD 0 3 7 1 20 18 20 238 18 17 20 20 18 17 19 17 18 17 22 255 NERO_QUADD_CFD 0 3 9 1 20 20 20 20 20 23 21 25 22 23 22 20 21 20 20 21

Fig.14:  An illustrative example of the files shaper_init_NERO.dat and cfd_init_NERO.dat.

To set the discriminator thresholds, we look at the ADC spectra. When the discriminator thresholds are high enough, the noise-peak in the ADC spectra disappears. Note that for this operation the ADC spectrum of each channel has to be filled in coincidence with the same discriminator channel. To do this is necessary to display the ADC spectra of each channel gating on the respective TDC channel.

If thresholds have not been set up, we strongly recommend looking at the spectra in logarithmic scale in order to see the real signal from the shapers.

# 5.2   ADC thresholds

The ADC thresholds are used to cut the ADC pedestals. They are set up in the file NEROtest/Standalone/readout/nerodaqini.tcl looking at the ADCs spectra when the ADCs inputs are disconnected. If the ADC thresholds have not been set up, we strongly recommend looking at the spectra in logarithmic.

5.3   Electronic checks

When NERO is running properly, preamplifier signals looks like are the ones shown in Fig.6). The most important constrain in defining the gates is to have the end of the TDC gate always before the computer finish processing data. As explained with reference to Fig.7, the TDC gate is provided by the Master Gate Live signal and stop by the timeout signal generated by the 200 (sec delayed Master Gate Live signal.

##### Usual operation in standalone mode

# 6.1   Background

Once the thresholds have been set up, the background of TDC, ADC and Scalers should be comparable. Fig.15 shows background rates measured with quadrant B, during a 10 hour run.

Fig.15: Left: scaler and TDC measured background. Rigtht: scaler and ADC measured background.

# 6.2   Efficiency

The Cf-252 spontaneous-fission neutron source (Z7153) at NSCL was used to measure the neutron detection efficiency. Because the source decays mainly by (-decay, the neutron emission rate has to be calculated by multiplying the actual activity of the source by the fission branching ratio (where the neutrons come from) and the number of neutrons emitted per fission. This operation is done automatically by using the Mathematica file ‘Activity_Cf252.nb’ in the directory NEROtest/Standalone/readout/. The file contains the characteristics of the source, taken from the literature, its composition determined on 12/1/1989 and its activity on 12/1/1989 as determined from the calibration done on 5/27/2007.

In order to correctly measure the neutron efficiency, the source has to be in the middle of NERO. This is achieved by using a special holder (see Fig.16).

Fig.16 the NERO source holder (left), placed in NERO while holding a source (right)

The NERO efficiency depends on the location of the source along the central hole; a maximum value of 31.4% was found in the exact center of NERO. Fig.17 shows the efficiency as a function of the horizontal position along NERO central hole.

Fig.17: NERO efficiency as a function of the position inside NERO.

##### PRINCIPLE OF NEUTRON EMISSION PROBABILITY MEASUREMENT

  - Delayed neutron-emission probability measurements

After a nucleus is implanted in the BCS, the emitted (-decay electron will provide the Master Gate. This signal will open a 200(sec gate that is used by NERO to detect the possible (-delayed neutrons. The width of this gate was obtained from the time needed by the neutrons to be moderated (around 150 (sec). The number of neutrons associated with the preceding (-decay, detected during the time gate, provides the basic information required to calculate the Pn-value of the corresponding nucleus. Fig.18 shows some recent Pn value measured with NERO.

Fig.18  Pn values measured with NERO:    124Ag, 120-122Pd, 117-120Rh.

NERO Manual

First version

June 2007

# Footnotes

NERO is divided in four quadrant A,B,C,D and the detectors are numbered according to quadrant (see Fig.4)

for other information about this module see http://jelinux.pico-systems.com/shapdisc.html

-0.8V and –1.8V for logic 0 and 1 levels, respectively

0V and –0.8V for logic 0 and 1 levels, respectively

This is because for every trigger, all the ADC channels are red; also those that have just noise.

# Text boxes

3He Proportional

Counters

BF3 Proportional

Counters

Polyethylene

Moderator

## Boron Carbide

Shielding

After -decay

Before -decay

Preamplifier

Metallic tab indicating the right thread-connection

Quadrant A and B stop signals

START GATE signal from latch 3

Quadrant C and D stop signals

Sum channel

Shaper test channel

Discriminator test channel

Logical OR between the 16-discriminator channels

Discriminator out

Shaper out

Preamplifier input